Six-point calibration method of VNA-O-Mizer by Daniel Marks KW4TI

The raw data output of the VNA is captured using the IMACQ command. The IMACQ command has three parameters:

IMACQ <# frequencies> <start frequency> <end frequency>

which captures the raw data for <# frequencies> from <start frequency> to <end frequency>. The result is in the following format, which is the raw data captured using quadrature demodulated samples:

<frequency>,<channel 1 in phase voltage>,<channel 1 quadrature voltage>, <channel 1 in phase current>,<channel 1 quadrature current>, <channel 2 in phase current>,<channel 2 quadrature current>

Three complex numbers are formed from this data (which are not actual voltage and current, but will be used to calibrate it):

V_1 = (\mbox{channel 1 in-phase voltage}) + j (\mbox{channel 1 quadrature voltage})

I_1 = (\mbox{channel 1 in-phase current}) + j (\mbox{channel 1 quadrature current})

I_2 = (\mbox{channel 2 in-phase current}) + j (\mbox{channel 2 quadrature current})

The number of averages which influences the dynamic range is set by

AVERAGES <# averages> <timeout in ms>

where the default number of averages is 64 for a 40 to 45 dB dynamic range, which increases to about 50 to 55 dB for 1000 averages which slows down the acquisition.

The calibration method is as follows:

For the one-port cal, capture the raw data for an open, short, and load (usually 50 ohms) fort channel 1. These will form six complex numbers V_{1}^{(open)} , I_{1}^{(open)} , V_{1}^{(short)} , I_{1}^{(short)} , V_{1}^{(load)} , I_{1}^{(load)} .

Divide these to form three trial impedances using complex division:

Z_1^{(open)} = \frac{V_1^{(open)}}{I_1^{(open)}} , Z_1^{(short)} = \frac{V_1^{(short)}}{I_1^{(short)}} , Z_1^{(load)} = \frac{V_1^{(load)}}{I_1^{(load)}}

When you actually measure an impedance at channel 1, you will measure a complex-valued voltage and current which is V_{1}^{(dut)} and I_{1}^{(dut)} . These are divided to form Z_{1}^{(dut)} = V_{1}^{(dut)} / I_{1}^{(dut)} . The actual impedance is then calibrated from the following bilinear transformation

Z_1^{(estimated)} = \frac{Z_1^{(dut)} + B }{C Z_1^{(dut)} + D }

where B, C, D, are the three complex-valued calibration parameters for port 1. These parameters are captured for each frequency. These are calculated using:

B = -Z_1^{(short)}

C = \frac{1}{Z_0} \frac{Z_1^{(short)} - Z_1^{(load)}}{Z_1^{(open)} - Z_1^{(load)}} where Z_0 is the characteristic impedance (usually 50 ohms)

D = -\frac{Z_1^{(open)}}{Z_0} \frac{Z_1^{(short)} - Z_1^{(load)}}{Z_1^{(open)} - Z_1^{(load)}} = - Z_1^{(open)} C

For port two calibration, first the isolation between port 1 and port 2 is found. The raw complex-valued current at port 2 is I_2 . I measure the current at port 2 during the port 1 short-circuit measurement step:

S_2 = \frac{I_2}{\mbox{\# averages}}

The total amount of signal leaking from port 1 to 2 scales with the number of averages, so the calibration is stored normalized by the number of averages.

To calibrate the impedance and gain at port 2, port 1 is connected to port 2. The raw current at port 2 is I_2 . The calibrated voltage and current at port 2 from port 1 are calculated from the raw samples V_1 and I_1 when they are connected together by

V_1^{(calib)} = V_1 + B I_1

I_1^{(calib)} = C V_1 + D I_1

and the following two parameters are captured:

Z_2 = \frac{V_1^{(calib)}}{I_2 - S_2 <\mbox{\# averages}>}

G_2 = \frac{I_1^{(calib)}}{I_2-S_2<\mbox{\# averages}>}

To calculate the complex-valued S_{11} , use

S_{11} = \frac{Z_1^{(estimated)} - Z_0}{Z_1^{(estimated)} + Z_0}

from which SWR, dB, and degrees may be determined.

To calculate S_{21}

S_{21} = \left(I_2 - S_2 <\mbox{\# averages}>\right) \frac{Z_2 + Z_0 G_2}{V_1^{(calib)} + Z_0 I_1^{(calib)}}

To calculate the thru impedance in series impedance mode

Z_{series} = \frac{V_1^{(calib)} - Z_2 \left(I_2 - S_2 <\mbox{\# averages}>\right)} {G_2 \left(I_2 - S_2 <\mbox{\# averages}>\right)}

and the thru impedance in shunt mode

Z_{shunt} = \frac{Z_2 \left(I_2 - S_2 <\mbox{\# averages}>\right)} {I_1^{(calib)} - G_2 \left(I_2 - S_2 <\mbox{\# averages}>\right)}